## NIST IR 6969

Selected Laboratory and Measurement Practices, and Procedures, to Support Basic Mass Calibrations

Georgia L. Harris

NIST Weights & Measures Division

Jose A. Torres

NIST Weights & Measures Division (contract)

This page is CODEN page on back of title page or Blank.

Foreword

This NIST IR of Selected Publications has been compiled as an interim update for a number of Good Laboratory Practices, Good Measurement Practices, Standard Operating Procedures, Statistical Techniques and Reference Tables.  This interim publication will be used as a reference document for the CD-ROM Mass Metrology Course being prepared by the NIST Weights and Measures Division.

Many of these procedures are updates to procedures that were originally published in NBS Handbook 145, Handbook for the Quality Assurance of Metrological Measurements, in 1986, by Henry V. Oppermann and John K. Taylor.   The updates have incorporated many of the requirements noted for procedures in ISO Guide 25, ANSI/NCSL Z 540-1-1994, and ISO/IEC 17025 laboratory quality systems. The major changes incorporate 1) uncertainty analyses that comply with current international methods and 2) measurement assurance techniques using check standards.  No substantive changes were made to core measurement processes or equations, with the exception of SOP 2, Standard Operating Procedure for Applying Air Buoyancy Corrections.  The CIPM 1981/91 equation for calculating air density has been added to SOP 2.

The following Practices and Procedures are new in this interim publication:

Good Measurement Practices for:

- Assignment and Adjustment of Calibration Intervals for Laboratory Standards (11)

- Standard Operating Procedure Selection (12)

- Ensuring Traceability (13)

- Selection and Use of Sensitivity Weights in Weighing Procedures (14)

Standard Operating Procedures for:

- Assignment of Uncertainty (29)

- Process Measurement Assurance Program (30)

The essential portions of SOP 29, Standard Operating Procedure for Assignment of Uncertainty were first published in 1994 for the NCSL Conference and Symposium and to assist State weights and measures laboratories comply with the ISO Guide to the Expression of Uncertainty in Measurements (1993).

Special thanks go to Val Miller and Jose Torres for the critical editorial reviews needed to complete this interim publication; to the Minnesota Metrology Laboratory for early drafts for GMP 11 and 13; and to Jerry Everhart for assistance with early drafts of SOP 30.

This page is blank.

Table of Contents

Foreword										iii

Good Laboratory Practices

| 1 | Quality Assurance of the Measurement Process | GLP 1-1 |
| --- | --- | --- |
| 4 | Periodic Recalibration Of State Standards | GLP 4-1 |
| 9 | Rounding Expanded Uncertainties and Calibration Values | GLP 9-1 |
| 11 | Painting Weights | GLP 11-1 |

Good Measurement Practices

| 1 | Reading Turning Points On An Equal Arm Balance | GMP 1-1 |
| --- | --- | --- |
| 10 | Understanding Factors Affecting Weighing Operations | GMP 10-1 |
| 11 | Assignment and Adjustment of Calibration Intervals for Laboratory Standards | GMP 11-1 |
| 12 | Standard Operating Procedure Selection | GMP 12-1 |
| 13 | Ensuring Traceability | GMP 13-1 |
| 14 | Selection and Use of Sensitivity Weights in Weighing Procedures | GMP 14-1 |

Standard Operating Procedures

| 1 | Preparation of Test/Calibration Reports | SOP 1-1 |
| --- | --- | --- |
| 2 | Applying Air Buoyancy Corrections | SOP 2-1 |
| 4 | Weighing by Double Substitution Using a Single Pan Mechanical Balance, a Full Electronic Balance, or a Balance with Digital Indications and Built In Weights | SOP 4-1 |
| 7 | Weighing by Single Substitution Using a Single Pan Mechanical Balance, a Full Electronic Balance, or a Balance with Digital Indications and Built In Weights | SOP 7-1 |
| 8 | Medium Accuracy Calibration of Mass Standards by Modified Substitution | SOP 8-1 |
| 9 | Control Charts for Calibration of Mass Standards | SOP 9-1 |
| 29 | Assignment of Uncertainty | SOP 29-1 |
| 30 | Process Measurement Assurance Program | SOP 30-1 |

Reference Materials

|  | Statistical Techniques (Updated Section 8 of NBS Handbook 145, 1986) | 8-1 |
| --- | --- | --- |
|  | Reference Tables (Updated Section 9 of NBS Handbook 145, 1986) | 9-1 |

# This page is blank.

# Good Laboratory Practices

# Good Measurement Practices

# Standard Operating Procedures

# Statistical Techniques and Reference Tables